## Slide 1: XPETS Module 2 Summary and User’s Manual
By Grant Buster
May 27, 2015

## Slide 2: XPETS Module 2 Calibration Geometry Solver

## Slide 3: Geometry Assumptions
Your system geometry is only defined by your x-ray point source, detector and turntable, NOT the test section
The detector is assumed to be a perfectly vertical plane
The detector is assumed not to be rotated in the xz plane (LEVEL IT!)
The origin is set as (0,0,0)
  The xy origin is set at the turntable rotational axis
  The z origin coordinate is set at the plane intersecting the x-ray point source and perpendicular to the detector plane (the z-origin plane is horizontal)
X-Ray Point Source
Turntable Rotation Axis
Set as Origin
Detector Plane

## Slide 4: Importance of Geometry
An accurate 3D reconstruction is impossible without a good geometry solution
What does this mean?
4 key geometric parameters need to be defined:
  Y_of (y axis distance from origin to focal/source point)
  O_x (x axis distance from right hand side of detector to origin)
  O_y (y axis distance from detector plane to origin)
  O_z (z axis distance from top of detector to origin)
X-Ray Point Source
Turntable Rotation Axis
Set as Origin
Detector Plane

## Slide 5: O_y and Y_of
O_y and Y_of have been defined using a machined reference geometry and Solidworks geometry solver
  O_y = 0.2283 (m)
  Y_of = 1.8453 (m)
These values should be considered constant unless:
  The turntable is moved
  The columnator or detector are moved in the y-direction (but vertical shifts should not affect O_y and Y_of)
SolidWorks Geo Solver (Top Down View, XY Plane)
Radius of Rotation
Reference Piece Diameter
Detector Plane
A)
B)
C)
A)
B)
C)
Reference Piece Diameter
Machined Reference Piece

## Slide 6: Defining O_x
We define O_x using an x-ray image set of a single point (usually a thumb screw) rotated through 360 degrees. NOTE: the screw should be placed as far off-center as possible (away from the center of the detector in z and x directions).
The angle and the i,j tip location (in pixels) (i=column#, j=row#) for each image need to be entered in the format:
  Ox_DATA(:,:,1)=[ [angle;nan] , [ j ; i ] ];
Into the code: XPREX_20141029_Full_Geo_Solver.m
Example:
XPREX_20141029_Full_Geo_Solver.m

## Slide 7: Defining O_x
XPREX_20141029_Full_Geo_Solver.m then utilizes the fact that the horizontal location of the screw tip will move in a sinusoidal path:
XPREX_20141029_Full_Geo_Solver.m will automatically calculate any angle offset and O_x
  Angle offset is caused by the screw not being placed perfectly at 0-degees when the turntable is at 0-degrees. This is accounted for and it is okay to place the screw at a random angle.

## Slide 8: Defining O_z
This is the (algebraically) difficult part
Start from a top-down view, in order to calculate the radius of rotation of the thumb screw tip
Radius of rotation is calculated by the following equation:
Where i is the horizontal pixel location (column index) of the thumbscrew tip in one image. This is usually taken when the screw is far right or far left. Theta in the above equation is theta of the turntable plus the theta offset
X-Ray Point Source
Turntable Rotation Axis
Set as Origin
Detector Plane
Thumb-Screw
Radius of Rotation

## Slide 9: Defining O_z
We move to a side-view, in order to calculate O_z
Two images are picked, usually the two where the thumbscrew tip is highest and lowest in the image
R1 and R2 are calculated using R_rot*sin(theta), such that:
O_z is calculated as follows:
X-Ray Point Source
Detector Plane
R2
R1
K1
K2
O_z

## Slide 10: Geometry Solved!
The geo.mat file should now be good to go!
Two other important parameters:
  The detector has 6993 pixels per meter (PPM)
  A pin/pebble diameter is 0.01257 (m)
The geo.mat file should be placed in the active directory, and loaded before module2 starts to run

## Slide 11: XPETS Module 2 3D Bed Reconstruction

## Slide 12: Module 2 Code Architecture
RUN_2015MMDD.m
onetwo3D_script.m
Place3D.m
Phase1_Parallel_Find.m
Phase2_Eliminate_Outsiders.m
Phase3_Eliminate_Overlaps.m
Validator.m
Render_Pebbles.m
Geo
Find2_W_Verify.m
Find_Contrast_1pin.m
TriangleIntersection.m
TopDownIntersection.m
Thresholds
Thresholds
Contrast Data
Geo
Output file

## Slide 13: RUN_2015MMDD.m
This file will execute the following:
  Setup parallel toolbox
  Iterate through study timesteps
    Import module1 data
    Take user selected threshold values
    Load geometry file
    Save output file
  Iterate through post processing
    Validator
    Render
This file is intended to hold all of the user inputs and setup parameters. High level organization.

## Slide 14: onetwo3D_script.m
This file will execute the following:
  (Optional) pre select a region of module1 data
    This is good for prototyping code
    Selects a small region of interest for faster runtime
  Place3D.m
  Phase1_Parallel_Find.m
  Phase2_Eliminate_Outsiders.m
  Phase3_Eliminate_Overlaps.m
This file is largely organizational in nature
See details on each phase in later slides

## Slide 15: Place3D.m
This file is responsible for taking module 1 data and placing the endpoints on their respective detector planes in 3-dimensional space
Highly geometry dependent
Input (module1) (in pixels):
Output (3D in meters):
Where point 1 and point 2 should be the endpoints of a pin in an x-ray image projected onto the detector plane in 3D space

## Slide 16: Place3D.m
Looks like this:

## Slide 17: Phase1_Parallel_Find.m
This phase sets up the parallel processing inputs in a way that MATLAB will accept
This phase iterates through possible comparisons of module 1 data, running the Find2_W_Verify.m algorithm
  For example, module 1 will return endpoints for 5 different detector angles
    The find2_w_verify algorithm will take data sets from 2 detector angles and compare them for possible 3D results
    Phase1_parallel_find.m iterates through possible input pairs for find2_w_verify.m
Phase1_Parallel_Find.m is organizational and is mostly meant to setup the parallel processing structure
Parallel processing can be easily swapped with serial by replacing the PARFOR loop with a FOR loop
  Everything else should be kept the same… setup is for parallel but can also work for serial, not the other way around

## Slide 18: Find2_W_Verify.m
This is the main reconstruction algorithm
The function takes in two sets of pairs of endpoints
The function iterates through both sets, trying all possible comparisons between two pairs of endpoints
An initial midpoint vector calculation is done to see if they are at all close vertically (easy way to eliminate possibilities)
If pairs of endpoints are close, run intersection algorithms
  TriangleIntersection.m
  TopDownIntersection.m
If intersection algorithms find a solution (within set thresholds) use a contrast verification algorithm to check solution against contrast data for all five images
If contrast data is verified, add solution to output

## Slide 19: TriangleIntersection.m
A endpoints belonging to the same pin in an image form a triangle with the third point being the x-ray point source
Vector calculus can be used to compare two triangles and find their linear intersection
A good vector calculus reference can be found here: http://geomalgorithms.com/a06-_intersect-2.html
The algorithm results in two line segment solutions:
  Endpoints1 as a triangle intersecting endpoints2 as a plane
  Endpoints2 as a triangle intersecting endpoints1 as a plane
Thresholds used to eliminate:
  Normalized length: are the linear solutions the same length as the pin’s true length (pebble diameter)?
  Metric spatial similarity: are the midpoints of the two linear solutions close together?
The two linear solutions will always be perfectly parallel

## Slide 20: TriangleIntersection.m
This is what a triangle solutions look like:
Endpoint Pair 2
Endpoint Pair 1
Triangle/Plane 2
Triangle/Plane 1
Possible solutions
Note: In this case, the solution is valid so the two possible line solutions overlap perfectly

## Slide 21: TopDownIntersection.m
As the pin approaches a horizontal orientation in the actual system, the triangle intersection method approaches massive uncertainty (think about the intersection of two near-horizontal planes)
The top down method addresses this issue by examining the intersection of the XY vector components, and then examining where these projected intersections actually lie in the 3-space
The thresholds are:
  Normalized length: are the linear solutions the same length as the pin’s true length (pebble diameter)?
  Metric spatial similarity: are there consistent solutions in 3-space?
This method returns more possible results than the triangle intersection method (see next slide)

## Slide 22: TopDownIntersection.m
The possible solutions look like this:
Top-down
Intersections
(possible solutions)
Actual
Solution
Top-Down View
Side View
Top-down
Intersections
Actual
Solution
Note: in this example, the pin is likely found by the triangle intersection method as well.

## Slide 23: Find_Contrast_1pin.m
Every possible solution is projected onto the x-ray contrast data for every detector angle (usually 5 angles) so that it can be verified that this is indeed a pin in each image
Vector math is used for the projections
  Similar to intersection algorithms
An average contrast threshold is set for each image angle (set in the run file)
5 pixels of “slop” is considered and factored into an quantitative uncertainty (in the output)
This function is called the most times, and is therefore the most computationally expensive function to run… It is, on the other hand, incredibly useful in making sure you actually found a solution

## Slide 24: Find_Contrast_1pin.m
The contrast projection looks like this:
0-degree
22.5-degree
45-degree
67.5-degree
90-degree
The output contrast for the 5 images:

## Slide 25: Find_Contrast_1pin.m
Detailed view:
Multiple contrast samples
account for uncertainty

## Slide 26: PHASE1 Summary
In summary, phase 1 looks like this:

## Slide 27: Phase2_Eliminate_Outsiders.m
This phase eliminates any possible solutions that may have been incorrectly identified as correct solutions outside of the physical silo
It uses max and min endpoints from module 1 to set silo boundaries in the 5 different views
It also uses two custom set silo boundaries at custom set angles (user input in the RUN.m file)
It only eliminates horizontal outliers using vertical boundaries

## Slide 28: Phase2_Eliminate_Outsiders.m
Looks like this:
Note: this image was created with very liberal thresholds and the result is therefore considerably more messy than it should be. Red lines are eliminated for being outside of the silo, blue lines are solutions, pink circles are module 1 results.

## Slide 29: Phase3_Eliminate_Overlaps.m
This module uses the physical constraint that pebbles should contact but not overlap
An overlap threshold is set (normalized to a pebble diameter) in the RUN.m file
If any two solutions overlap more than the threshold, one is eliminated based on the uncertainty set forth by the find contrast function

## Slide 30: Phase3_Eliminate_Overlaps.m
Before Overlap Elimination
After Overlap Elimination

## Slide 31: Validator.m
This simply projects the solutions onto the x-ray image so the user can visually verify the results
Various options can be found in the code for labels, image saving, module1 plotting and other “accessories”

## Slide 32: Render_Pebbles.m
This function plots pebbles with the option to plot the pebble equator, color and transparency
Very helpful for post-processing modules

## Slide 33: XPETS Module 2 Summary

## Slide 34: User Input Summary
Inputs found in the RUN.m file:
  TriangleIntersection Thresholds
    Normalized Length
      0 to 1, how close to the true pin length is the possible solution? 1 is exactly the correct length
    Metric Spatial
      In meters. How close are the possible solutions in 3D metric coordinates? Perfectly close together (threshold=0) is a better solution.
  TopDownIntersection Thresholds (also uses normalized length)
    Similar Z (metric)
      In meters. How close are possible solutions in the z-direction? Perfectly close together (threshold=0) is a better solution.

## Slide 35: User Input Summary
Inputs found in the RUN.m file:
  Elimination Thresholds
    Silo Boundary (Theta, Max, Min)
      If you were to take an x-ray image at the specified theta (degrees), the min and the max values (in pixels) specify left and right silo boundaries. Any solutions found to be left of min and right of max will be eliminated
    Contrast Threshold
      Average contrast of solution projected onto the image must be greater than the contrast threshold
    Number of Valid images
      The solution must have contrast greater than the contrast threshold in the specified number of images. 5 signifies that the solution must be validated in all 5 images
    Normalized overlap
      Fraction of a pebble diameter that two pins can overlap. 1 means two pebbles cannot overlap at all (centroids must be 1 diameter apart). 0.5 means that centroids can be 1 radius apart and not be eliminated.

## Slide 36: Output Summary
Module 2 saves output files with outputs from each phase (1,2,3) as well as the related diagnostics files. Phase3 is the final output
The output data structure is as follows:
Where the 1 and 2 signify the two endpoints of a physical pin in 3-dimensions for n number of pins
And the corresponding diagnostics data structure is:
Diagnostics definitions on the next slide

## Slide 37: Output Diagnostics Format Details
Length, colinearity and spatial similarity are all threshold-type variables set by the intersection algorithms
P is the 3D module 1 point (from Place3D.m) that the solution originated from
S is the 3D source point corresponding to P
C is the average contrast of the solution in each image (1-5)
U is an uncertainty parameter set by the find contrast algorithm

## Slide 38: DON’T PANIC Always feel free to call 720-495-6245 or email grantbuster@gmail.com